## Sheet1
Bidirectional Traceability Matrix
AI-Based Lung Opacity Detection System · IEC 62304:2006 + A1:2015 · Version 1.0 · Bhavya Reddy
Test IDTest | FR ID | Requirement | Pipeline Level | Module | IEC Class | Verification Type | Test Description | Test Steps | Expected Result | Actual Result | Status
TC-01 | FR-01 | Image Input | Preprocessing | main.py | Class B | Manual | Verify PNG and JPG formats are accepted | Upload PNG → Analyze. Upload JPG → Analyze via frontend | Both formats accepted without error | PNG and JPG accepted via Pillow RGB conversion | Pass
TC-02 | FR-02 | Binary Classification | Level 1 | main.py | Class B | Automated | Verify Level 1 model classifies normal and abnormal X-rays correctly | python3 -m pytest scripts/test_verification.py::TestLevel1 -v -s | Normal-15.png → NORMAL, Lung_Opacity-2839.png → ABNORMAL | NORMAL (100.0% confidence), ABNORMAL (100.0% confidence) | Pass
TC-03 | FR-03 | Disease Classification | Level 2 | main.py | Class B | Manual | Verify Level 2 classifies disease into correct category | Upload COVID-19, opacity, and pneumonia X-rays and check Level 2 result | Correct disease category returned for each image | 98.0% overall. Per-class: Normal 100%, COVID-19 98%, Opacity 96%, Pneumonia 98% | Pass
TC-04 | FR-04 | Disease Probability Scores | Level 2 | main.py | Class B | Manual | Verify 4 probability scores are returned and sum to 100% | Upload abnormal X-ray and check Level 2 probability bar chart | 4 probability scores summing to 100% | YOLOv8 softmax output guarantees sum = 100% | Pass
TC-05 | FR-05 | Opacity Detection | Level 3 | main.py | Class B | Manual | Verify opacity regions are detected and reported with bounding boxes | Upload opacity X-ray → verify boxes detected. Upload normal X-ray → verify no boxes | Boxes on abnormal images, none on normal images | mAP50 = 53.1% — dataset quality ceiling documented | Pass
TC-06 | FR-06 | Lung Laterality Determination | Level 3 | main.py | Class B | Manual | Verify left, right, and bilateral lung involvement is correctly identified | Upload left, right, and bilateral opacity X-rays and check location field | Correct laterality returned for each case | cx < 0.5 → Left lung, cx > 0.5 → Right lung. Bilateral verified. | Pass
TC-07 | FR-07 | Affected Area Calculation | Level 4 + Segmentation | main.py | Class B | Manual | Verify affected area percentage is returned in valid range | Upload opacity X-ray and verify Level 4 value is between 0 and 100 | Numeric percentage between 0.0 and 100.0 | Intersection formula verified. Returns valid range. | Pass
TC-08 | FR-08 | Severity Classification | Level 5 | main.py | Class B | Automated | Verify all severity thresholds and bilateral upgrade rule | python3 -m pytest scripts/test_verification.py::TestSeverity -v -s | MILD (20%), MODERATE (35%), SEVERE (65%), MODERATE upgrade (20% bilateral) | All 4 tests passed. Bilateral upgrade rule verified. | Pass
TC-09 | FR-09 | Clinical Recommendation | Level 6 | main.py | Class B | Manual | Verify correct priority, action, and timeline for each severity level | Upload SEVERE, MODERATE, MILD X-rays and check Level 6 recommendation | IMMEDIATE (RED) 30 min, HIGH (ORANGE) 2-4 hrs, ROUTINE (YELLOW) 24 hrs | All 3 priority levels generated correctly with correct timelines | Pass
TC-10 | FR-10 | Analysis Output | Full Pipeline | main.py | Class B | Manual | Verify all 6 level results are returned within 10 seconds | Upload X-ray, start timer, verify all 6 fields in response within 10 seconds | Complete JSON response with all 6 fields within 10 seconds | Pipeline completes within 10 seconds. All 6 fields present in JSON. | Pass
*TC-05 / FR-05: Passes functionally. mAP50 = 53.1% is below 70% target — dataset quality ceiling. 5 retraining iterations conducted. Radiologist always reviews before clinical action.